Instructions for Upgrading Echoport USB Firmware FOR V1 BOARDS

Determine the current USB firmware version.

This first section details how to determine the current firmware version if required.

Check to see a firmware label is present.

(This is placed on the base of the Echoport USB, next to the serial number)

Or

Use the ‘System Check’ feature in EZ-Screen/V6 to read the firmware version.

Connect the Echoport USB to the PC using the USB cable.

Open EZ-Screen/V6.

From the main menu, select File>System Checks…

On the window that appears select Run Diagnostics.

Select ‘No’ to running hardware diagnostics.

Scroll up to the top of the window.

Read the USB Firmware version.

Close the window and close EZ-Screen/V6.

Preparing the hardware for programming.

Disconnect the USB cable from the back of the Echoport USB device.

Using a pointed device (i.e. a biro) move the position of the switch on the rear of the Echoport USB device from Normal to Service.

Connect the USB cable to the USB device.

Running the Update program.

Place the software CD into the computers disc drive and wait for the menu to appear.

Select browse CD.

Open folder V1 Firmware Programmer.

Double click on UsbOaeUpdate.exe

Click on the ‘Update Now!’ button.

Wait for the firmware to be updated, which could take a few minutes.

The message ‘All software updates complete!’ will appear.

Click OK.

Close the UsbOaeUpdate.exe program.

Reset the hardware for general use.

Disconnect the USB cable from the USB device.

Using a small pointed device move the position of the switch on the rear of the Echoport USB device from Service to Normal.